## Cover Letter
MEMORANDUM | (LABORATORY DATA REPORT)
EPA - General Parameters
|  |  |  | In reply refer to: | 22-MS-34
To: |  | Ken Forshay |  | From: | M. Sexton
|  |  |  | Lab: | General Parameters 123
Thru: |  | Mark White |  | Date: | 2022-06-21
Kristie Rue
Technical Directive No.: |  | EPAGPA076 |  | Originator: | Ken Forshay
Task No.: |  | AE2.6.2 |  | Copies: | Ken Forshay
Jake Beaulieu
Katherine Buckler
Kristie Rue
Molly Sexton
Anna Bachmann
Sample Site/Project: |  | Lake Methane Project
Date Collected: |  | 2022-06-13 | Sample Set No.: |  | 7868
Date Received: |  | 2022-06-15 |  | Sample Matrix: | water
Date Analyzed: |  | 2022-06-15 |  | Analysis Type: | NPOC & NPDOC
No. Samples Analyzed: |  | 2 filtered and unfiltered |  | Sample Preparation: | none
K-GCRD-SOP-1165-0 Determination of Various Fractions of Carbon in Aqueous Samples using the Shimadzu TOC-VCPH Analyzer.
Method(s) Used :
Comments:
The samples were analyzed for NPOC using the Shimadzu TOC-L Analyzer. Quality control measures were performed according to the criteria in K-GCRD-1165-0 for NPOC. All reagents, standards, dilutions, and lab blanks were prepared using Barnstead Nanopure Water from Lab 123. An MDL study for NPOC range 0-10 mg/L was performed on 10/25/2021. The current MDL for this range is 0.049 mg/L.

## Data
EPA - General Parameters
Laboratory Data Report
Laboratory: | General Parameters 123
Technical Directive: | EPAGPA076 |  |  | Sample Results
Analyst: | M. Sexton | Analytes | NPOC |  |  | NPDOC
|  | Codes | 7440-44-0-NPOC |  |  | 7440-44-0-NPDOC
Methods: | See Cover Letter | Methods | K-GCRD-SOP-1165-0 |  |  | K-GCRD-SOP-1165-0
|  | Unit | mg/L |  |  | mg/L
|  | MDL | 0.049 |  |  | 0.049
|  | QL | 0.5 |  |  | 0.5
Field Sample ID | Lab Sample ID | Date Collected | Date Analyzed | Data | DF | Date Analyzed | Data | DF
184DU | 7868-1 | 2022-06-13 | 2022-06-15 | 8.15 | 1 | 2022-06-15 | 8.23 | 1
184SU | 7868-2 | 2022-06-13 | 2022-06-15 | 8.37 | 1 | 2022-06-15 | 7.98 | 1
184SU | 7868-2 LAB DUP | 2022-06-13 | 2022-06-15 | 8.15 (RPD = 2.66) | 1 | 2022-06-15 | 7.92 (RPD = 0.755) | 1
Comments: The data quality objective for the precision of sample duplicates is a relative percent difference (RPD) of < 10%, which was met for this set of samples, that are within the calibration range. An MDL study for NPOC high sensitivity range 0-10mg/L was performed on 10/25/2021.
Notes:
1. If the parameter was detected above the quantitation limit (QL), the numeric result is reported; BQL denotes that the parameter was not detected at or above the quantitation limit; BQL ( ) denotes that the parameter was detected above the method detection limit (MDL) but below QL and the estimated numeric result is reported in parenthesis; ND denotes that the parameter was not detected at all. All the results are corrected with dilution factors (DF), if applicable. NA means not applicable.
2. " -" denotes that the information is not available or the analyte is not analyzed.

## QC Data
EPA - General Parameters
Laboratory Data Report
Laboratory: | General Parameters 123
Tech. Directive: | EPAGPA076 |  | Quality Control Data Summary
Analyst: | M. Sexton |  | Analytes | NPOC |  |  | NPDOC
|  |  | Codes | 7440-44-0-NPOC |  |  | 7440-44-0-NPDOC
|  |  | Methods | K-GCRD-SOP-1165-0 |  |  | K-GCRD-SOP-1165-0
|  |  | Unit | mg/L |  |  | mg/L
|  |  | MDL | 0.049 |  |  | 0.049
|  |  | QL | 0.5 |  |  | 0.5
QC Sample ID | Additional ID | Date Prepared | Date Analyzed | Data | True Value | % REC. |  | True Value | % REC.
MB | Blank Nanopure | 2022-06-15 | 2022-06-15 | ND | - | - | ND | - | -
MB | Blank Nanopure | 2022-06-15 | 2022-06-15 | ND | - | - | ND | - | -
SS | ERA Cat. 669 | 2022-05-23 | 2022-06-15 | 1.68 | 1.58 | 106 | 1.68 | 1.58 | 106
SS | ERA Cat. 669 | 2022-05-23 | 2022-06-15 | 1.72 | 1.58 | 109 | 1.72 | 1.58 | 109
CCC | Calibration Check Standard | 2022-06-15 | 2022-06-15 | BQL (0.491) | 0.5 | 98.2 | BQL (0.491) | 0.5 | 98.2
CCC | Calibration Check Standard | 2022-06-15 | 2022-06-15 | 9.85 | 10 | 98.5 | 9.85 | 10 | 98.5
CCC | Calibration Check Standard | 2022-06-15 | 2022-06-15 | 1.02 | 1 | 102 | 1.02 | 1 | 102
CCC | Calibration Check Standard | 2022-06-15 | 2022-06-15 | 4.82 | 5 | 96.4 | 4.82 | 5 | 96.4
CS | Caffeine Standard | 2022-06-15 | 2022-06-15 | 4.96 | 5 | 99.2 | 4.96 | 5 | 99.2
CS | Caffeine Standard | 2022-06-15 | 2022-06-15 | 4.96 | 5 | 99.2 | 4.96 | 5 | 99.2
MS | 184DU Spike | 2022-06-15 | 2022-06-15 | 13.5 | 8.15 (4.98) | 107 | 13.5 | 8.23 (4.98) | 106
Comments:
The data quality objective for the accuracy of continuing calibration check standards and second source standards for NPOC is 90-110% recovery. The data quality objective for the accuracy of matrix spikes is 80-120% recovery. The data quality objective for the accuracy of the caffeine standard is ≥70% recovery. These objectives were met for the standards and matrix spikes. The matrix spikes for NPOC and NPDOC were prepared by adding 100 µL of a 1000 mg/L standard into 20 mL of sample yielding a spike concentration of 4.98 mg/L. The matrix spike recovery was calculated according to the equation: % Recovery = 100 x (Spiked sample concentration (DATA) - Native sample concentration)/Spike concentration.
Notes:
1. MB - Method Blank. CCC - Continuing Calibration Check. A calibration standard analyzed within the batch of samples. CS- A Caffeine Standard. A caffeine standard analyzed as an evaluation of the instrument. LCS - Laboratory Control Spike. A laboratory blank spiked with analytes at known concentrations. MS - Matrix Spike. A field sample spiked with known concentrations of analytes. The field sample id is identified. The True Value column for matrix spikes list the unspiked native sample concentration along with the spike concentration in parentheses. SS - Samples obtained from the second sources are identified by their designated names. DUP - Field sample duplicate analysis. A sample selected by the lab analyst to analyze as a duplicate. It is reported in the sample result section. % REC - Percent Recovery. Calculated as the percentage of the results to the true values. It equals to % accuracy for CCC.