## Cover Letter
MEMORANDUM | (LABORATORY DATA REPORT)
EPA - General Parameters
|  |  |  | In reply refer to: | 22-MS-57
To: |  | Ken Forshay |  | From: | Molly Sexton
|  |  |  | Lab: | General Parameters
Thru: |  | Mark White |  | Date: | 2022-09-19
Kristie Rue
Technical Directive No.: |  | EPAGPA081 |  | Originator: | Ken Forshay
RAP No.: |  | AE 2.6.2 |  | Copies: | Ken Forshay
Jake Beaulieu
Katy Buckler
Kristie Rue
Molly Sexton
Anna Bachmann
Sample Site/Project: |  | Lake Methane Project
Date Collected: |  | 9/12 & 9/15/2022 | Sample Set No.: |  | 7904
Date Received: |  | 2022-09-16 |  | Sample Matrix: | water
Date Analyzed: |  | 2022-09-19 |  | Analysis Type: | NO2 + NO3, NH4,
No. Samples Analyzed: |  | 6 |  | Sample Preparation: | See comments
Method(s) Used : |  | Ammonia Lachat FIA 10-107-06-1-O
Nitrate+Nitrite Lachat FIA 10-107-04-2-A
K-GCRD-SOP-1151-1 Quality Control Procedures for General Parameters Analysis
Using Lachat Flow Injection Analyzers.
Comments:
The quality control measures performed are outlined in K-GCRD-SOP-1151-1. The MDL's were determined on 9/8/2022 for NO3+NO2 and NH4. The current MDL's are 0.021 mg/L for NO2+NO3 and 0.032 mg/L for NH4.

## Data
EPA - General Parameters
Laboratory Data Report
Laboratory: | General Parameters
Technical Directive: | EPAGPA081 |  |  | Sample Results
Analyst: | Molly Sexton | Analytes | Nitrate+Nitrite-N (NO3+NO2-N) |  |  | Ammonia-N(NH4-N)
|  | Codes | 630 |  |  | 6484-52-2
|  | Methods | K-GCRD-SOP-1151-1 |  |  | K-GCRD-SOP-1151-1
|  | Unit | mg/L |  |  | mg/L
|  | MDL | 0.021 |  |  | 0.032
|  | QL | 0.1 |  |  | 0.1
Field Sample ID | Lab Sample ID | Date Collected | Date Analyzed | Data | DF | Date Analyzed | Data | DF
011DU | 7904-1 | 2022-09-12 | 2022-09-19 | ND | 1 | 2022-09-19 | 0.816 | 1
011NB | 7904-2 | 2022-09-12 | 2022-09-19 | ND | 1 | 2022-09-19 | ND | 1
011SD | 7904-3 | 2022-09-12 | 2022-09-19 | ND | 1 | 2022-09-19 | BQL (0.033) | 1
011SU | 7904-4 | 2022-09-12 | 2022-09-19 | BQL (0.025) | 1 | 2022-09-19 | ND | 1
003DU | 7904-5 | 2022-09-15 | 2022-09-19 | BQL (0.029) | 1 | 2022-09-19 | 0.207 | 1
003SU | 7904-6 | 2022-09-15 | 2022-09-19 | BQL (0.035) | 1 | 2022-09-19 | BQL (0.052) | 1
003SU | 7904-6 LAB DUP | 2022-09-15 | 2022-09-19 | BQL (0.027) (RPD=NA) | 1 | 2022-09-19 | BQL (0.037) (RPD=NA) | 1
BHP 532 DUP |  |  | 2008-08-08
Comments:
The data quality objective for the precision of sample duplicates is a relative percent difference (RPD) of < 10%, which was met for this set of samples, that are within the calibration range. Results reported as BQL are numeric estimates (See Notes below). The MDL's were determined on 9/8/2022 for NO3+NO2 and NH4.
Notes:
1. If the parameter was detected above the quantitation limit (QL), the numeric result is reported; BQL denotes that the parameter was not detected at or above the quantitation limit; BQL ( ) denotes that the parameter was detected above the method detection limit (MDL) but below QL and the estimated numeric result is reported in parenthesis; ND denotes that the parameter was not detected at all. All the results are corrected with dilution factors (DF), if applicable. NA means not applicable.
2. " -" denotes that the information is not available or the analyte is not analyzed.

## QC Data
EPA - General Parameters
Laboratory Data Report
Laboratory: | General Parameters
Tech. Directive: | EPAGPA081 |  | Quality Control Data Summary
Analyst: | Molly Sexton |  | Analytes | Nitrate+Nitrite-Nitrogen (NO3+NO2-N) |  |  | Ammonia-N(NH4-N)
|  |  | Codes | 630 |  |  | 6484-52-2
|  |  | Methods | K-GCRD-SOP-1151-1 |  |  | K-GCRD-SOP-1151-1
|  |  | Unit | mg/L |  |  | mg/L
|  |  | MDL | 0.021 |  |  | 0.032
|  |  | QL | 0.1 |  |  | 0.1
QC Sample ID | Additional ID | Date Prepared | Date Analyzed | Data | True Value | % REC. | Data | True Value | % REC.
MB | Blank Nanopure | 2022-09-19 | 2022-09-19 | ND | - | - | ND | - | -
MB | Blank Nanopure | 2022-09-19 | 2022-09-19 | ND | - | - | ND | - | -
SS | ERA # 72 | 2018-01-18 | 2022-09-19 | 3.65 | 3.47 | 105 | 12.1 | 11.4 | 106
SS | ERA # 72 | 2018-01-18 | 2022-09-19 | 3.54 | 3.47 | 102 | 12.1 | 11.4 | 106
CCC | Calibration Check Standard | 2022-08-16 | 2022-09-19 | 0.103 | 0.1 | 103 | BQL (0.094) | 0.1 | 94
CCC | Calibration Check Standard | 2022-08-16 | 2022-09-19 | 18.9 | 20 | 94.5 | 20 | 20 | 100
CCC | Calibration Check Standard | 2022-08-16 | 2022-09-19 | 0.961 | 1 | 96.1 | 0.985 | 1 | 98.5
CCC | Calibration Check Standard | 2022-08-16 | 2022-09-19 | 9.32 | 10 | 93.2 | 9.93 | 10 | 99.3
MS | 011DU Spike | 2022-09-19 | 2022-09-19 | 6.27 | ND (6.17) | 102 | 6.98 | 0.816 (6.17) | 99.9
Comments:
The data quality objective for the accuracy of continuing calibration check standards and second source standards is 90-110% recovery. The data quality objective for the accuracy of matrix spikes is 80-120% recovery. These objectives were met for the standards and matrix spikes. The matrix spikes for NO3+NO2 and NH4 were prepared by adding 50 µL of a 500 mg/L standard into 4 mL of sample yielding a spike concentration of 6.17 mg/L. The matrix spike recovery was calculated according to the equation: % Recovery = 100 x (Spiked sample concentration (DATA) - Native sample concentration)/Spike concentration.
Notes:
1. MB - Method Blank. CCC - Continuing Calibration Check. A calibration standard analyzed within the batch of samples. LCS - Laboratory Control Spike. A laboratory blank spiked with analytes at known concentrations. MS - Matrix Spike. A field sample spiked with known concentrations of analytes. The field sample id is identified. The True Value column for matrix spikes list the unspiked native sample concentration along with the spike concentration in parentheses. SS - Samples obtained from the second sources are identified by their designated names. DUP - Field sample duplicate analysis. A sample selected by the lab analyst to analyze as a duplicate. It is reported in the sample result section. % REC - Percent Recovery. Calculated as the percentage of the results to the true values. It equals to % accuracy for CCC.